## Slide 1
Co-assembly of Zwitterionic and Cationic Polymers
for Antibacterial and Antithrombotic Surfaces
via Weak Electrostatic Interactions
Wang et al. | Biomaterials 331 (2026) 124114
LbL coating · Zwitterionic polymers · Antibacterial · Antithrombotic
PA2.5
Optimized dual-function coating
>97.6%
Bactericidal efficiency vs S. aureus
3.1%
Thrombus coverage (vs PLA 99.0%)

## Slide 2
Table of Contents
Wang et al. | Biomaterials 331 (2026) 124114
2
01
Background & Motivation
Clinical need for dual-function medical device coatings
02
Materials & Methods
Polymer synthesis, LbL coating, characterization toolkit
03
Solution Interactions
ITC, turbidity, fluorescence — weak co-assembly confirmed
04
Coating Properties
QCM-d, AFM, XPS, WCA, zeta potential
05
Antifouling & Antithrombotic
Ex vivo blood contact — PA2.5: 3.1% thrombus coverage
06
Antibacterial Performance
Contact sterilization — PA2.5: >97.6% bactericidal
07
Biocompatibility & In Vivo
Hemolysis, cytotox, rat/rabbit animal models

## Slide 3
Wang et al. | Biomaterials 331 (2026) 124114
3
SECTION 01
Background & Motivation
Clinical need for dual-function medical device coatings

## Slide 4
01 · BACKGROUND
The Dual Challenge: Infection and Thrombosis
Wang et al. | Biomaterials 331 (2026) 124114
4
Bacterial Infection
Medical device-associated infections account for ~2 million cases per year in the USA alone. Biofilm formation on implant surfaces leads to antibiotic resistance and implant failure.
Thrombosis
Thrombus formation on implanted materials triggers device failure, embolic events, and systemic complications. Platelet adhesion and activation occur within seconds of blood contact.
Antibiotic Resistance
Conventional antibiotic coatings provide only short-term protection and accelerate resistance. Contact-killing surfaces offer a sustained, resistance-free antibacterial mechanism.
Conflicting Strategies
Cationic surfaces kill bacteria but promote thrombosis. Zwitterionic surfaces are antifouling but lack intrinsic antibacterial activity. Achieving both simultaneously remains a key challenge.

## Slide 5
01 · BACKGROUND
Current Approaches and Their Limitations
Wang et al. | Biomaterials 331 (2026) 124114
5
Zwitterionic Coatings
·
Excellent antifouling & antithrombotic via hydration layer
·
No inherent antibacterial activity
·
Polysulfobetaine, polycarboxybetaine well-studied
Cationic (QA) Coatings
·
Membrane disruption kills bacteria on contact
·
Cationic charge promotes platelet activation
·
Antithrombotic performance is poor
Antibiotic-Releasing Coatings
·
Short-term protection only
·
Drives antibiotic resistance
·
Drug depletion limits long-term use
Dual-Function Attempts
·
Sequential layers sacrifice one function
·
Blending disrupts both functional groups
·
No prior success via co-assembly approach

## Slide 6
01 · BACKGROUND
Design Concept: Weak Co-assembly for Dual Function
Wang et al. | Biomaterials 331 (2026) 124114
6
Key Insight: When zwitterionic and cationic polymers co-assemble via weak electrostatic interactions, both functional groups remain accessible — enabling simultaneous antifouling and antibacterial activity from a single coating layer.
PA (Poly-carboxybetaine amide)
Zwitterionic polymer
COO⁻ and N⁺ groups
Strong hydration layer → antifouling
Weak interaction with PQ
PS (Poly-sulfobetaine)
Zwitterionic polymer
SO₃⁻ and N⁺ groups
Excellent antifouling properties
Different interaction mode with PQ
PQ (Poly-quaternary ammonium)
Cationic polymer
Permanent positive charge
Membrane disruption → antibacterial
Accessible when interaction is weak
Co-assembly Result
PA-PQ or PS-PQ complex
LbL coating on PLA substrate
Retains both functional groups
Dual antibacterial + antithrombotic

## Slide 7
01 · BACKGROUND
Research Overview
Wang et al. | Biomaterials 331 (2026) 124114
7
Hypothesis
Weak electrostatic interaction between zwitterionic (PA/PS) and cationic (PQ) polymers preserves functional group accessibility, enabling dual-function LbL coatings.
Solution-Phase Characterization
ITC, fluorescence spectroscopy, turbidimetry, and viscometry confirm spontaneous co-assembly (ΔG < 0) with weak interaction for PA-PQ pair.
Surface Characterization
QCM-d, AFM, XPS, WCA, and zeta potential quantify coating deposition, morphology, elemental composition, and hydrophilicity as a function of formulation.
In Vitro Performance
Ex vivo blood contact tests, MIC assays, and CFU counting evaluate antithrombotic and antibacterial performance. PA2.5 achieves 3.1% thrombus coverage and >97.6% kill.
In Vivo Validation
Rat subcutaneous, rabbit jugular vein, and PGLA suture implant models confirm safety and efficacy. PA2.5 achieves >99% bacterial eradication in vivo.

## Slide 8
Wang et al. | Biomaterials 331 (2026) 124114
8
SECTION 02
Materials & Methods
Polymer synthesis, LbL coating process, and characterization toolkit

## Slide 9
02 · METHODS
Polymer Components: Structure and Properties
Wang et al. | Biomaterials 331 (2026) 124114
9
PA — Poly(carboxybetaine amide)
[Figure Placeholder]
Chemical structure
·
Zwitterionic: pendant COO⁻ and N⁺ groups
·
Amphoteric character enables weak interaction with PQ
·
Strong hydration layer at physiological pH
·
Synthesized via RAFT polymerization
·
Confirmed by ¹H NMR and GPC
PS — Poly(sulfobetaine)
[Figure Placeholder]
Chemical structure
·
Zwitterionic: pendant SO₃⁻ and N⁺ groups
·
Balanced dipole — no net charge
·
Known for protein and cell resistance
·
Synthesized via RAFT polymerization
·
Confirmed by ¹H NMR and GPC
PQ — Poly(quaternary ammonium)
[Figure Placeholder]
Chemical structure
·
Permanently cationic at all pH values
·
Membrane-disrupting antibacterial activity
·
Interacts electrostatically with PA or PS
·
Interaction strength tunable by ratio
·
Confirmed by ¹H NMR and GPC

## Slide 10
02 · METHODS
Layer-by-Layer Coating Process & Sample Nomenclature
Wang et al. | Biomaterials 331 (2026) 124114
10
1
Surface
Preparation
PLA film cleaned with ethanol and DI water
2
Co-assembly
Solution
Mix zwitterionic + PQ at defined concentrations
3
Dip Coating
Substrate immersed in co-assembly solution
4
Drying &
Curing
Air-dried; process repeated for multilayers
5
Characterization
QCM-d, AFM, XPS, WCA, antibacterial/antithrombotic
Sample Nomenclature
Sample
Zwitterionic polymer
PQ conc. (mg/mL)
Key property
PA2.0
PA — 2 mg/mL
0.5
Antifouling-dominant
PA2.5
PA — 2 mg/mL
1.0
Optimized dual function ★
PS2.0
PS — 2 mg/mL
0.5
Antibacterial (PQ accessible)
PS2.5
PS — 2.5 mg/mL
1.0
Reduced efficacy (PQ shielded)

## Slide 11
02 · METHODS
Characterization Toolkit
Wang et al. | Biomaterials 331 (2026) 124114
11
QCM-d
Quartz Crystal Microbalance
with Dissipation
Real-time mass deposition
Film stiffness (ΔD)
Assembly kinetics
AFM
Atomic Force
Microscopy
Surface morphology
Roughness: Rq, Ra
Nano-scale topography
XPS
X-ray Photoelectron
Spectroscopy
Elemental composition
Surface chemistry
N, S, O quantification
WCA
Water Contact
Angle
Surface hydrophilicity
PA2.5: 21.58° (hydrophilic)
PLA: 88.73° (hydrophobic)
ITC
Isothermal Titration
Calorimetry
Thermodynamics of co-assembly
ΔH, ΔS, ΔG, Ka
Interaction strength
Fluorescence
Fluorescence
Spectroscopy
Complex tightness
PS-PQ: avg 34.25 (tight)
PA-PQ: avg 43.59 (loose)

## Slide 12
Wang et al. | Biomaterials 331 (2026) 124114
12
SECTION 03
Solution Interactions
ITC, fluorescence, turbidity, and viscometry — characterizing the co-assembly

## Slide 13
03 · SOLUTION INTERACTIONS
ITC and Turbidity: Thermodynamics of Co-assembly
Wang et al. | Biomaterials 331 (2026) 124114
13
Isothermal Titration Calorimetry (ITC)
[Figure Placeholder]
ITC thermogram — PA-PQ and PS-PQ
ΔG < 0
Both pairs: thermodynamically spontaneous co-assembly
PA-PQ
Weak electrostatic interaction — carboxylate groups
PS-PQ
Relatively stronger hydrophobic/electrostatic interaction
Key difference
PA-PQ weak binding → PQ accessible for antibacterial action
Turbidity & Aggregate Formation
[Figure Placeholder]
Turbidity vs concentration — aggregate size
·
PA-PQ: turbidity increases at higher concentrations → larger aggregates
·
PS-PQ: more compact aggregates (tighter complex)
·
Loose PA-PQ assembly → functional groups remain exposed on surface
·
Concentration ratio determines coating properties and performance

## Slide 14
03 · SOLUTION INTERACTIONS
Fluorescence and Viscosity: Complex Tightness
Wang et al. | Biomaterials 331 (2026) 124114
14
[Figure Placeholder]
Fluorescence intensity vs titration ratio — PA-PQ vs PS-PQ
[Figure Placeholder]
Viscosity — individual vs co-assembly solutions
34.25
PS-PQ avg intensity
Tighter complex
43.59
PA-PQ avg intensity
Looser complex
Mechanistic Interpretation
PS-PQ forms a tighter complex (lower fluorescence emission, smaller aggregates). PA-PQ forms a looser complex — the carboxylate groups of PA interact weakly with PQ's quaternary ammonium. This weaker interaction is key: it allows PQ functional groups to remain accessible at the coating surface for antibacterial activity, while PA groups maintain the hydration layer for antifouling.

## Slide 15
Wang et al. | Biomaterials 331 (2026) 124114
15
SECTION 04
Coating Properties
QCM-d, AFM, XPS, zeta potential, and water contact angle

## Slide 16
04 · COATING PROPERTIES
QCM-d: Real-Time Coating Deposition
Wang et al. | Biomaterials 331 (2026) 124114
16
[Figure Placeholder]
QCM-d frequency (Δf) and dissipation (ΔD) vs time — PA-PQ and PS-PQ
PA-PQ
Δf = 156 Hz
ΔD = 559 × 10⁻⁶
More mass deposited
Highly hydrated soft film → superior antifouling
PS-PQ
Δf = 89 Hz
ΔD = 225 × 10⁻⁶
Less mass, stiffer film
Compact, denser coating

## Slide 17
04 · COATING PROPERTIES
Zeta Potential: Surface Charge Analysis
Wang et al. | Biomaterials 331 (2026) 124114
17
[Figure Placeholder]
Zeta potential after each coating layer — PLA, PS-PQ, PA-PQ series
Key Findings
Bare PLA substrate: negative zeta potential
After PS-PQ coating: shifts toward near-neutral charge balance
After PA-PQ coating: slightly negative — favorable for antithrombotic surfaces
PA2.5 surface charge reduces electrostatic attraction with blood proteins
Charge evolution consistent with layer-by-layer buildup confirmed by QCM-d

## Slide 18
04 · COATING PROPERTIES
AFM: Surface Morphology and Roughness
Wang et al. | Biomaterials 331 (2026) 124114
18
[Figure Placeholder]
AFM height images — PLA, PA-PQ (layers 1–5), PS-PQ (layers 1–5)
[Figure Placeholder]
Rq and Ra vs number of coating layers — PA-PQ vs PS-PQ
PLA (bare)
Rq = 1.06 nm
Ra = 0.67 nm
Baseline roughness
PA-PQ
Rq decreases
Smoother surface
Each layer fills nanoscale defects → uniform hydration layer
PS-PQ
Rq increases
Rougher surface
Compact aggregates create higher topographic variation

## Slide 19
04 · COATING PROPERTIES
Water Contact Angle: Surface Hydrophilicity
Wang et al. | Biomaterials 331 (2026) 124114
19
[Figure Placeholder]
WCA photographs and values — PLA, PA2.0, PA2.5, PS2.0, PS2.5
PLA (control)
88.73°
Moderately hydrophobic
PA2.0
53.3°
Hydrophilic (PA coating)
PA2.5
21.58°
Highly hydrophilic ★
PS series
>60°
More hydrophobic than PA
Mechanism: The zwitterionic PA groups form a dense hydration shell through hydrogen bonding and electrostatic interaction with water molecules. Higher PQ content (PA2.5) increases surface charge density, further enhancing hydrophilicity.

## Slide 20
Wang et al. | Biomaterials 331 (2026) 124114
20
SECTION 05
Antifouling & Antithrombotic Performance
Ex vivo whole blood contact — protein, platelet, and thrombus evaluation

## Slide 21
05 · PERFORMANCE
Protein Adsorption and Platelet Adhesion
Wang et al. | Biomaterials 331 (2026) 124114
21
[Figure Placeholder]
Fluorescence images — protein adsorption (BSA-FITC) on PLA vs PA2.5
[Figure Placeholder]
SEM images — platelet adhesion and activation on PLA vs coated surfaces
Antifouling Performance
Protein adsorption
PA2.5 dramatically reduces BSA adsorption vs bare PLA. The hydration layer acts as a steric/energetic barrier.
Platelet adhesion
Significantly fewer platelets adhere to PA2.5 surfaces. Adhered platelets show rounded (inactive) morphology vs activated on PLA.
Mechanism
Highly hydrated PA groups (WCA 21.58°) repel hydrophobic proteins and inhibit platelet integrin engagement with the surface.
PA vs PS
PA2.5 outperforms PS series due to greater hydrophilicity and more uniform coating morphology (smoother AFM topography).

## Slide 22
05 · PERFORMANCE
Ex Vivo Antithrombotic: Whole Blood Contact Test
Wang et al. | Biomaterials 331 (2026) 124114
22
[Figure Placeholder]
Photographs and SEM of tube surfaces after whole-blood contact — PLA, PS, PA series
Thrombus Coverage (%)
PLA
99.0%
PS2.0
97.1%
PS2.5
89.7%
PA2.0
30.3%
PA2.5
3.1%
Thrombus Mass Increase (%)
PLA
152.9%
PS2.0
90.3%
PS2.5
82.2%
PA2.0
52.5%
PA2.5
51.7%

## Slide 23
Wang et al. | Biomaterials 331 (2026) 124114
23
SECTION 06
Antibacterial Performance
MIC assay, solution bactericidal efficiency, and contact sterilization

## Slide 24
06 · ANTIBACTERIAL
MIC Assay and Solution Bactericidal Efficiency
Wang et al. | Biomaterials 331 (2026) 124114
24
[Figure Placeholder]
MIC heatmap — PA/PQ concentration ratios vs S. aureus / E. coli growth
[Figure Placeholder]
Colony plates — solution bactericidal test PA-PQ vs PS-PQ
MIC Results (PA/PQ mg/mL)
PA 2 / PQ 0.5
Effective
PA 2 / PQ 1.0
Effective
PA 2 / PQ 2.0
Effective
PA 2 / PQ 3.0
Effective
PS-PQ — Formulation-Dependent Efficacy
PS2.0
97.41%
Effective — PQ accessible
PS5.0
97.04%
Effective — PQ accessible
PS2.5
<10%
Ineffective — PQ shielded by excess PS
PS5.5
<10%
Ineffective — PQ shielded by excess PS

## Slide 25
06 · ANTIBACTERIAL
Contact Sterilization: In Vitro Antibacterial Surface
Wang et al. | Biomaterials 331 (2026) 124114
25
[Figure Placeholder]
Colony-forming unit (CFU) plates — PLA, PA2.0, PA2.5, PS2.0, PS2.5
6.87×10³
CFU/cm²
Bare PLA
~0
CFU/cm²
PA2.0 coating
~0
CFU/cm²
PA2.5 coating ★
Mechanism: Contact-Killing
·
PQ quaternary ammonium groups remain surface-accessible on PA2.5
·
Electrostatic attraction draws negatively-charged bacterial membrane
·
Membrane disruption → intracellular contents leak → cell death
·
PA zwitterionic groups prevent non-specific adhesion of debris
·
No antibiotic released → no resistance pressure
·
PA2.0 and PA2.5 both achieve near-complete (~100%) surface sterilization

## Slide 26
Wang et al. | Biomaterials 331 (2026) 124114
26
PA2.5 — OPTIMIZED DUAL-FUNCTION COATING
Key Performance Results
3.1%
Thrombus coverage
(vs PLA 99.0%)
>97.6%
Bactericidal efficiency
vs S. aureus
21.58°
Water contact angle
(highly hydrophilic)
~0 CFU
CFU/cm² on surface
(contact sterilization)
>99%
In vivo bacterial
eradication (rat model)
<5%
Hemolysis rate
(non-hemolytic)
Why PA2.5?
✓
Weak PA-PQ electrostatic interaction preserves PQ accessibility
✓
Highly hydrated surface (WCA 21.58°) repels blood proteins
✓
Smooth morphology (AFM) enables uniform functional coverage
✓
Dual function confirmed in 3 animal models
✓
Biocompatible: hemolysis <5%, L929 viability >80%

## Slide 27
Wang et al. | Biomaterials 331 (2026) 124114
27
SECTION 07
Biocompatibility & In Vivo Validation
Hemolysis, cytotoxicity, and three animal model studies

## Slide 28
07 · BIOCOMPATIBILITY
Hemolysis and Cytotoxicity: In Vitro Safety
Wang et al. | Biomaterials 331 (2026) 124114
28
[Figure Placeholder]
Hemolysis rate (%) — PLA, PA2.0, PA2.5, PS2.0, PS2.5
[Figure Placeholder]
L929 cell viability (%) — MTT assay — all formulations at 24/48/72 h
<5%
Hemolysis rate
All groups non-hemolytic
(< 5% threshold)
>80%
L929 cell viability
All groups biocompatible
at all time points

## Slide 29
07 · BIOCOMPATIBILITY
In Vivo: Rat Subcutaneous Implant Model
Wang et al. | Biomaterials 331 (2026) 124114
29
[Figure Placeholder]
H&E staining and immunofluorescence (CD3, CD68) — day 1 and day 3 post-implant
[Figure Placeholder]
CFU recovered from tissue — PA2.5 vs control day 1 and day 3
>99%
Bacterial eradication
at day 1
PA2.5 implant
>99%
Bacterial eradication
at day 3
PA2.5 implant
Immune Response
·
H&E: PA2.5 shows mild tissue reaction vs significant inflammation in control
·
CD3+ T-cell infiltration: significantly reduced around PA2.5 implant
·
CD68+ macrophage infiltration: reduced — low chronic inflammation
·
PA2.5 demonstrates favorable host immune response profile

## Slide 30
07 · BIOCOMPATIBILITY
In Vivo: Rabbit Jugular Vein Antithrombotic Model
Wang et al. | Biomaterials 331 (2026) 124114
30
[Figure Placeholder]
Photographs and histology of jugular vein sections — PLA, PS2.5, PA2.5 after implantation
PLA (Control)
Significant thrombus formation
Lumen markedly narrowed
Platelet aggregation extensive
PS2.5 Coating
Substantial thrombus
Some improvement vs PLA
Cationic charge limits antifouling
PA2.5 Coating
Minimal thrombus
Lumen largely patent
Confirms ex vivo results in vivo ★

## Slide 31
07 · BIOCOMPATIBILITY
In Vivo: PGLA Suture Antibacterial Model (Rat)
Wang et al. | Biomaterials 331 (2026) 124114
31
[Figure Placeholder]
CFU from tissue around PGLA sutures — PA2.5-coated vs uncoated — day 1 and day 3
[Figure Placeholder]
Fluorescence images of bacteria around suture wound — day 1 and day 3
PGLA Suture Model — Key Findings
>99%
Bacterial eradication around PA2.5-coated suture at day 1
>99%
Bacterial eradication around PA2.5-coated suture at day 3
·
Demonstrates practical applicability: biodegradable suture material
·
Coating persists through early wound healing period
·
Reduced immune cell infiltration vs uncoated suture control
·
Supports translation to clinical suture and implant applications

## Slide 32
Wang et al. | Biomaterials 331 (2026) 124114
32
SECTION 08
Conclusions
Mechanism, comparative summary, and future outlook

## Slide 33
08 · CONCLUSIONS
Proposed Mechanism: Weak Co-assembly Enables Dual Function
Wang et al. | Biomaterials 331 (2026) 124114
33
[Figure Placeholder]
Schematic: weak PA-PQ interaction → exposed PQ (antibacterial) + hydrated PA (antifouling)
Co-assembly
PA and PQ co-assemble spontaneously (ΔG < 0) via weak electrostatic interactions between carboxylate and quaternary ammonium groups.
Coating formation
LbL deposition creates a stable, smooth, hydrated film on PLA. PA2.5 produces the most uniform and hydrophilic coating (WCA 21.58°).
Antifouling/Antithrombotic
Dense hydration layer from PA zwitterionic groups repels proteins and platelets. Low surface energy prevents thrombus initiation (3.1% coverage).
Antibacterial
Weak interaction leaves PQ functional groups surface-accessible. Membrane disruption kills bacteria on contact (>97.6% efficiency) without antibiotic release.
Biocompatibility
Zwitterionic character minimizes non-specific interactions. Hemolysis <5%, cell viability >80% at all tested concentrations.

## Slide 34
08 · CONCLUSIONS
Performance Comparison: PA2.5 vs Controls
Wang et al. | Biomaterials 331 (2026) 124114
34
Metric
PLA (bare)
PS2.0
PS2.5
PA2.0
PA2.5 ★
Thrombus coverage
99.0%
97.1%
89.7%
30.3%
3.1%
Thrombus mass incr.
152.9%
90.3%
82.2%
52.5%
51.7%
WCA (°)
88.73
~70
~65
53.3
21.58
Solution bactericidal
—
97.41%
<10%
>97%
>97.6%
Surface CFU/cm²
6870
~high
—
~0
~0
Hemolysis
<5%
<5%
<5%
<5%
<5%
Cell viability
>80%
>80%
>80%
>80%
>80%
In vivo bacterial kill
—
—
—
—
>99%

## Slide 35
Wang et al. | Biomaterials 331 (2026) 124114
35
CONCLUSIONS & IMPACT
Conclusions and Future Outlook
First demonstration
PA-PQ co-assembly via weak electrostatic interactions achieves genuine dual antibacterial and antithrombotic function from a single coating system.
Key quantitative outcomes
PA2.5 coating: thrombus coverage 3.1% (vs PLA 99.0%), bactericidal efficiency >97.6%, in vivo bacterial eradication >99%, hemolysis <5%.
Mechanism validated
The weak PA-PQ interaction is essential: strong interaction (as in PS-PQ) shields PQ groups and compromises antibacterial activity. Weak interaction preserves both functions.
Translational potential
Demonstrated on PLA films and PGLA sutures. Applicable to catheters, orthopedic implants, cardiovascular devices. LbL process is substrate-agnostic and scalable.
Future Directions
→
Long-term stability and durability studies under physiological conditions
→
Clinical translation: catheter and cardiovascular device coating trials
→
Expanded pathogen spectrum: Gram-negative bacteria, fungi, viruses
→
Biodegradable formulations for transient implant applications